## Command Comparision
Category | Description |  |  | 4294A | E4990A
IEEE Common Commands |  |  |  | *CLS | *CLS
|  |  |  | *ESE | *ESE
|  |  |  | *ESR | *ESR
|  |  |  | *IDN? | *IDN?
|  |  |  | *OPC | *OPC
|  |  |  | *OPC? | *OPC?
|  |  |  | *OPT? | *OPT?
|  |  |  | *RST | *RST
|  |  |  | *SRE | *SRE
|  |  |  | *STB | *STB
|  |  |  | *TRG | *TRG :TRIG:IMM
|  |  |  | *WAI | *WAI
Meas. Condition | Preset |  |  | *RST | *RST
|  |  |  | PRES | :SYST:PRES
| Active Trace |  |  | TRAC | :CALC:PAR:SEL
| Meas Para |  |  | MEAS | :CALC:PAR:DEF
| Number of points |  |  | POIN | :SENS:SWE:POIN
| Sweep | Parameter |  | SWPP | :SENS:SWE:TYPE
|  | Type |  | SWPT
|  | Direction |  | SWED | :SENS:SWE:DIR
|  | Time |  | SWET | :SENS:SWE:TIME
|  | Delay | Sweep | SDELT | :SENS:SWE:DEL
|  |  | Meas. Point | PDELT | :SENS:SWE:PDEL
|  | Range | Start | STAR | :SENS:FREQ:STAR :SOUR:VOLT:STAR :SOUR:CURR:STAR :SOUR:BIAS:VOLT:STAR :SOUR:BIAS:CURR:STAR
|  |  | Stop | STOP | :SENS:FREQ:STOP :SOUR:VOLT:STOP :SOUR:CURR:STOP :SOUR:BIAS:VOLT:STOP :SOUR:BIAS:CURR:STOP
|  |  | Center | CENT | :SENS:FREQ:CENT :SOUR:VOLT:CENT :SOUR:CURR:CENT :SOUR:BIAS:VOLT:CENT :SOUR:BIAS:CURR:CENT
|  |  | Span | SPAN | :SENS:FREQ:SPAN :SOUR:VOLT:SPAN :SOUR:CURR:SPAN :SOUR:BIAS:VOLT:SPAN :SOUR:BIAS:CURR:SPAN
|  | Manual Sweep | On/Off | MANS | Not Available
|  |  | Specified Meas. Point | MANP | Not Available
|  |  | Specified Range | MANR | Not Available
| Signal Source（OSC) | OutputLevel | Setup Mode | POWMOD | :SOUR:MODE
|  |  | Setting | POWE | :SOUR:VOLT :SOUR:CURR
|  | Frequency |  | CWFREQ | :SENS:FREQ
| DC Bias | On/off |  | DCO | :SOUR:BIAS:STAT
|  | Range |  | DCRNG | :SOUR:BIAS:RANG
|  | Mode |  | DCMOD | :SOUR:BIAS:MODE
|  | Level | Current | DCI | :SOUR:BIAS:CURR
|  |  | Voltage | DCV | :SOUR:BIAS:VOLT
|  |  | Voltage | MAXDCV | :SOUR:BIAS:VOLT:LIM:MAX
|  |  | Min.Voltage | MINDCV | :SOUR:BIAS:VOLT:LIM:MIN
| Bandwidth |  |  | BWFACT | :SENS:APER
| Averaging | Per sweep | On/off |  | :CALC:AVER:STAT
|  |  | Factor | AVERFACT | :CALC:AVER:COUN
|  |  | Restart | AVERREST | :CALC:AVER:CLE
|  | Meas. Per point | On/off |  | :SENS:AVER:STAT
|  |  | Factor | PAVERFACT | :SENS:AVER:COUN
List Sweep | Span display type (Single/Segment) |  |  | LISPAN | :DISP:WIND:X:SPAC
| List | Edit |  | EDITLIST | Use the segment sweep
|  | Clear |  | CLEL | :SENS:SWE:TYPE SEGM
|  | Done |  | EDITDONE | :SENS:SEGM:DATA
| Segment Creation | Edit |  | SEDI
|  | Add |  | SADD
|  | Determine target segment |  | SEGM
|  | Delete |  | SDEL
|  | Done |  | SDON
|  | Trace color |  | SCOL | Not Available
Display | On/Off |  |  | SCRN | :DISP:ENAB
| Backlight On/Off |  |  | BLIGHT | :SYST:BACK
| Trace A/B split display ON/Off |  |  | SPLD | :DISP:WIND:SPLit
| Displayed trace setting |  |  | DISP | :DISP:WIND:TRAC:STAT :DISP:WIND:TRAC:MEM
| Copy from Meas.Data to Memory |  |  | DATMEM | :CALC:MATH:MEM
| Accumulate display On/Off |  |  | ACCUD | :DISPl:WIND:TRAC:ACC
| Non-active trace display On/Off |  |  | HIDI | Not Available
| Allocation (Measurement/IBASIC) setting |  |  | DISA | Not Available
| Voltage/Current Monitor Display Setup | Signal Source（OSC) |  | OMON | AC Level mesurement trace :CALC:PAR:DEF VAC or IAC, then get the data of trace data.
|  | DC Bias |  | BMON | DC measurement trace :CALC:PAR:DEF VDC or IDC, then get the data of trace data. DC Level Monitor :SENS:DC:MEAS:ENAB
| Title Setup |  |  | TITL | :DISP:WIND:TITL:DATA :DISP:WIND:TITL
| Offset Value Setup |  |  | DATOVAL | :CALC:TRAC:MATH:OFFS
| Display Format |  |  | FMT | :DISP:WIND:TRAC:Y:SPAC :DISPl:WIND:TRAC:GRAT:FORM
| ２D graph Display Format |  |  | x | :CALC:FORM:TWOD
| Unit Setup at PhaseDisplay |  |  | PHAU | :CALC:TRAC:UNIT:ANGL
| Extended PhaseDisplay On/Off at Phase Display |  |  | EXPP | :CALC:TRAC:FORM
| Display Scale | Auto Scale |  | AUTO | :DISP:WIND:TRAC:Y:AUTO
|  | Grid （Linear/LogYAxis Format） | Min. Value | BOTV | :DISP:WIND:TRAC:Y:BOTT
|  |  | Max. Value | TOPV | :DISP:WIND:TRAC:Y:TOP
|  |  | Reference line Position | REFP | :DISP:WIND:TRAC:Y:RPOS
|  |  | Reference line Value | REFV | :DISP:WIND:TRAC:Y:RLEV
|  |  | Scale/Div Value | SCAL | :DISP:WIND:TRAC:Y:PDIV
|  | Grid （Complex Format） | Xaxis Reference Value | REFX | :DISP:WIND:TRAC:X:RLEV
|  |  | Yaxis Reference Value | REFY | :DISP:WIND:TRAC:Y:RLEV
|  |  | Scale/Div Value | SCAL | :DISP:WIND:TRAC:Y:PDIV
|  | Polar Format Full Scale Value |  | REFV | :DISP:WIND:TRAC:Y:PDIV
|  | Scaled trace (Data/Memory) | Coupled scale On/Off | SCAC | Not Available
|  |  | Trace selection | SCAF | Not Available
Limit Test | On/Off |  |  | LIMITEST | :CALC:TRAC:LIM:STAT
| Beep On/Off |  |  | BEEPFAIL | :SYST:BEEP:WARN:STAT
| Table Creation | Edit Start |  | EDITLIML | :CALC:LIM:DATA
|  | Clear |  | LIMCLEL | This command define the limit test table.
|  | Edit Done |  | LIMEDONE
| Segment Creation | Edit Start |  | LIMSEDI
|  | Add |  | LIMSADD
|  | Determine target segment |  | LIMSEGM
|  | Delete |  | LIMSDEL
|  | Edit |  | LIMSQUI
|  | Edit Done |  | LIMSDON
|  | Parameter Setup | Start | LIMSTAR
|  |  | Stop | LIMSTOP
|  |  | Start Upper Value | LIMUSTAR
|  |  | Start Lower Value | LIMLSTAR
|  |  | Start Middle Value | LIMVSTAR
|  |  | Start Delta | LIMDSTAR
|  |  | Stop Upper Value | LIMUSTOP
|  |  | Stop Lower Value | LIMLSTOP
|  |  | Stop Middle Value | LIMVSTOP
|  |  | Stop Delta | LIMDSTOP
|  |  | On/Off | LIMSTEST
| Limit Range Offset | Limit value (X-axis) offset |  | LIMIAMPO | :CALC:TRAC:LIM:OFFS:STIM
|  | Sweep parameter value (Y-axis) offset |  | LIMIPRMO | :CALC:TRAC:LIM:OFFS:AMPL
| Test ResultOutput | All measurement points |  | OUTPLIML? | :CALC:LIM:REP:ALL
|  | Marker Position |  | OUTPLIMM? | Not Available
|  | Faied Measurement Point |  | OUTPLIMF? | :CALC:LIM:REP:DATA
|  | Number of Failed Measurment Point |  | OUTPFAIP? | :CALC:LIM:REP:POIN
Adapter Setup | Adapter Selection |  |  | E4TP | :SENSE:ADAP:TYPE :SENSE:FIXT:SEL
| Data for setup data calculation | Meas. |  | ECAL{P¦A¦B¦C} | :SENSE:ADAPter:CORR:COLL:ACQ
|  | Calculate and store to EEPROM |  | ECALDON | :SENSE:ADAPter:CORR:COLL:SAVE
|  | Meas.Abort |  | ECALQUI | :SENSE:ADAPter:CORR:COLL:CLE
User Cal | Standard for data measurement | Definition | Open | DCALOPEN{G¦C} | :SENS:CORR1:CKIT:OPEN:{G¦C}
|  |  | Short | DCALSHOR{R¦L} | :SENS:CORR1:CKIT:SHORt:{R¦L}
|  |  | Load | DCALLOAD{R¦L} | :SENS:CORR1:CKIT:LOAD:{L¦R}
|  | Display Defined value |  | CALS | Not Available (Standard values are displayed as the softkey label)
| On/Off |  |  | CALST | :SENS:CORR1:STAT
| Data measurement point setting (User/Fixed) |  |  | CALP | :SENS:CORR1:COLL:FPO
| Data for calibration coefficient calculation | Meas. |  | CAL{A¦B¦C} | :SENS:CORR1:COLL
|  | Calculate and store to RAM |  | CALDON | :SENS:CORR1:COLL:SAVE
|  | Measurement cancel |  | CALQUI | :SENS:CORR1:COLL:CLE
|  | Input data |  | INPUCALC{1-3} | Not Available
|  | Output data |  | OUTPCALC{1-3}? | Not Available
Port extension compensation | On/Off |  |  | PORE | :SENS:CORR2:EDEL:STAT
| Port extension value | Delay Time |  | PORTZ | :SENS:CORR2:EDEL:TIME
|  | Electrical length |  | PORTL | :SENS:CORR2:EDEL:LENG
Fixture compensation | Standard for data measurement | Definition | Open | DCOMOPEN{G¦C} | :SENS:CORR2:CKIT:OPEN:[C¦G]
|  |  | Short | DCOMSHOR{R¦L} | :SENS:CORR2:CKIT:SHORt:[R¦L]
|  |  | Load | DCOMLOAD{R¦L} | :SENS:CORR2:CKIT:LOAD:[L¦R]
|  | Screen ListDisplay |  | COMS | Not Available (Standard values are displayed as the softkey label)
| On/Off |  |  | COMST{A¦B¦C} | :SENSE:CORR2:OPEN :SENSE:CORR2:SHOR :SENSE:CORR2:LOAD
| Data measurement point setting (User/Fixed) |  |  | CALP | :SENSE:CORR1:COLL:FPO
| Data for compensation coefficient calculation | Meas. |  | COM{A¦B¦C} | :SENSE:CORR2:COLL
|  | Input data |  | INPUCOMPC{1-3} | Not Available
|  | Output data |  | OUTPCOMC{1-3}? | Not Available
Meas. | Sweep | Mode Setting | Continuous | CONT | :INIT:CONT
|  |  | Hold | HOLD | :INIT:CONT
|  | Operation | Single | SING | :INIT:IMM
|  |  | Specified Number | NUMG | :TRIG:AVER
| Trigger | Source Selection |  | TRGS | :TRIG:SOUR
|  | Polarity of external trigger signal |  | TRGP | :TRIG:EXT:SLOP
|  | Trigger event mode |  | TRGEVE | :TRIG:POIN
|  | GPIB/LAN triggering |  | *TRG | *TRG: :TRIG:IMM
Data read/write | Data transfer format selection | ASCII Format |  | FORM4 | :FORMat:DATA ASC
|  | IEEE 32-bit floating point format |  | FORM2 | :FORMat:DATA REAL32
|  | IEEE 64-bit floating point format |  | FORM3 | :FORMat:DATA REAL
|  | MS-DOS Format |  | FORM5 | Not Available
| Query | DataArray | All Meas. Point | OUTPDATA? | :CALC:DATA:SDAT
|  |  | Specified Meas. Point | OUTPDATAP? | Not Available
|  | Memory | All Meas. Point | OUTPMEMO? | :CALC:DATA:SMEM
|  |  | Specified Meas. Point | OUTPMEMOP? | Not Available
|  | Data TraceArray | All Meas. Point | OUTPDTRC? | :CALC:DATA:FDAT
|  |  | Specified Meas. Point | OUTPDTRCP? | Not Available
|  | Memory TraceArray | All Meas. Point | OUTPMTRC? | :CALC:DATA:FMEM
|  |  | Specified Meas. Point | OUTPMTRCP? | Not Available
|  | SweepParameterArray | All Meas. Point | OUTPSWPRM? | :CALC:DATA:XAX
|  |  | Specified Meas. Point | OUTPSWPRMP? | Not Available
|  | Level MonitorSignal Source（OSC)Voltage | All Meas. Point | OUTPVAC? | Not Available
|  |  | Specified Meas. Point | OUTPVACP? | Not Available
|  | Level MonitorSignal Source（OSC)Current Value | All Meas. Point | OUTPIAC? | Not Available
|  |  | Specified Meas. Point | OUTPIACP? | Not Available
|  | Level MonitorDC Bias | All Meas. Point | OUTPDC? | DC Bias measurement trace :CALC:PAR:DEF VDC or IDC, then get the data of trace data.
|  |  | Specified Meas. Point | OUTPDCP? | Not Available
|  |  | Ch先頭 | X | :SENS:DC:MEAS:DATA：DCV?
|  |  | Ch先頭 | X | :SENS:DC:MEAS:DATA：DCI?
|  | Occurred error information (Error queue) |  | OUTPERRO? | :SYST:ERR
| Write | DataArray |  | INPUDATA | :CALC:DATA:SDAT
|  | DataTraceArray |  | INPUDTRC | :CALC:DATA:FDAT
Marker | Marker movement setting | Movement mode (Continuous/Discrete) |  | MKRMOV | :CALC:MARK:DISC
|  | Trace A/B coupled mode On/Off |  | MKRCOUP | :CALC:MARK:COUP
| Objective trace selection (Data/Memory) |  |  | MKRO | Not Available
| Marker (Main marker) | On/Off |  | MKR | :CALC:MARK
|  | Read Market Value | Meas.Parameter | MKRVAL? | :CALC:MARK:Y
|  |  | 2nd Meas.Parameter | MKRAUV? | :CALC:MARK:Y
|  |  | SweepParameter | MKRPRM | :CALC:MARK:X
|  |  | Meas./Sweep parameter value | OUTPMKR? | Use :CALC:MARK:X and :CALC:MARK:Y
|  | Move | Specified Meas. Point | MKRP | :CALC:MARK:X:POS
|  |  | SweepParameter | MKRPRM | :CALC:MARK:X
| Sub marker | On/Off |  | SMKR{1-7} | :CALC:MARK
|  | Read Market Value | Meas.Parameter | SMKRVAL{1-7}? | The E4990A has no sub marker but 9 markers.
|  |  | 2nd Meas.Parameter | SMKRAUV{1-7}? | Use the markers.
|  |  | SweepParameter | SMKRPRM{1-7}
|  |  | Meas./Sweep parameter value | OUTPSMKR{1-7}?
|  | Move | Specified Meas. Point | SMKRP{1-7}
|  |  | SweepParameter | SMKRPRM{1-7}
| Delta Marker | Mode Setting |  | DMKR | :CALC:MARK:REF :CALC:MARK:FUNC:TRAC
|  | Read Market Value | Meas.Parameter Value | DMKRVAL | :CALC:MARK:Y
|  |  | SweepParameter | DMKRPRM | :CALC:MARK:X
|  |  | SweepParameter | OUTPDMKR? | Use :CALC:MARK:X and :CALC:MARK:Y
|  | Move | Specified Meas. Point | DMKRP | :CALC:MARK:X:POS
|  |  | Meas.Parameter Value | DMKRVAL | Not Available (E4990A does not have the delta marker)
|  |  | 2nd. Meas.Parameter | DMKRAUV | Not Available
|  |  | SweepParameter | DMKRPRM | :CALC:MARK:X
| Unit of sweep parameter value |  |  | MKRXUNIT | :CALC:MARK:X:UNIT
| Level monitor value display mode setting |  |  | MKRMON | :CALC:MARK:MON
| Marker list function On/Off |  |  | MKRL | :DISP:TABL:TYPE MARK :DISP:TABL:STAT
| Marker sweep parameter value | Marker Position -> Start |  | MKRSTAR | :CALC:MARK:SET STAR
|  | Marker Position -> Stop |  | MKRSTOP | :CALC:MARK:SET STOP
|  | Marker Position -> Center | Set | MKRCNT | :CALC:MARK:SET CENT
|  |  | Set and Zoom Span | MKRZM | :CALC:MARK:SET ZOOM
|  |  | Zooming Aperture Setting | ZMAPER | CALC:MARK:ZAP
|  | Measurement point of manual sweep |  | MKRMANP | Not Available
|  | Limit Range | -> Start | MKRLIMSTAR | Not Available
|  |  | -> Stop | MKRLIMSTOP | Not Available
| Difference of sweep parameter between marker and delta marker -> Span |  |  | MKRDSPAN | Not Available
| MarkerPosition Meas.Parameter Value | -> Grid reference value |  | MKRREF | :CALC:MARK:SET RLEV
|  | Limit Range | -> X-axis offset value of limit range | MKRAMPO | Not Available
|  |  | -> Start point middle value of limit range | MKRLIMVSTAR | Not Available
|  |  | -> Stop point middle value of limit range | MKRLIMVSTOP | Not Available
Marker Search | Search Range Setting | Partial search On/Off |  | PARS | :CALC:MARK:FUNC:DOM
|  | Upper Limit | Specified Meas. Point | SEARMAXP | Not Available
|  |  | Specified SweepParameter Value | SEARMAX | :CALC:MARK:FUNC:DOM:STOP
|  |  | Marker position | MKRTRMAX | Not Available
|  | Lower Limit | Specified Meas. Point | SEARMINP | Not Available
|  |  | Specified SweepParameter Value | SEARMIN | :CALC:MARK:FUNC:DOM:STAR
|  |  | Marker position | MKRTRMIN | Not Available
|  | Range between marker and deleta marker |  | MKRTR | Not Available
| Search range setting for List sweep | Search range (All segments/Specified segment) |  | SEARNG | Use Search range
|  | Search segment |  | SEGMNUM | :CALC:MARK:FUNC:DOM
|  | Search range line display On/Off |  | SEARLINE
| Search tracking function On/Off |  |  | TRACK | :CALC:MARK:FUNC:TRAC
| Peak definition | Delta X Value Setting |  | PKDLTX | Not Available
|  | Delta Y Value Setting |  | PKDLTY | :CALC:MARK:FUNC:PEXC
|  | Polarity setting |  | PKPOL | :CALC:MARK:FUNC:PPOL
|  | Automatic definition using marker |  | MKRPKD | Not Available
| Search (Maximum value/Minimum value/Peak/Target value) |  |  | SEAM | :CALC:FUNC:TYPE {MAX¦MIN¦PEAK} :CALC:FUNC:EXEC
| Peak Search | Next peak search |  | SEANPK | Not Available
|  | Left Peak search |  | SEANPKL | :CALC:MARK:FUNC:TYPE LPE
|  | Right Peak search |  | SEANPKR | :CALC:MARK:FUNC:TYPE RPE
|  | Peak value -> sweep center value |  | PEAKCENT | Not Available
| Target Search | Target value setting |  | SEATARG | :CALC:MARK:FUNC:TARG
|  | Target line display On/Off |  | SEATARGL | Not Available
|  | Left target search |  | SEAL | :CALC:MARK:FUNC:TYPE LTAR
|  | Right target search |  | SEAR | :CALC:MARK:FUNC:TYPE RTAR
Meas. Result Analysis | Equivalent Circuit Analysis | Equivalent CircuitModel | Selection | EQUC | :CALC:EPAR:CIRC
|  |  | Display On/Off | DISECIRC | :CALC:EPAR:DISP
|  | Equivalent circuit parameter calculation |  | CALECPARA | :CALC:EPAR
|  | Equivalent CircuitParameter | Display On/Off | DISECPARA | :CALC:EPAR:DISP
|  |  | Read/Definition | DEFEC{R1¦C1¦L1¦C0} | :CALC:EPAR:CIRC:[A¦B¦C¦D¦E}:[R1¦L1¦C1¦R0]
|  | Frequency response simulation |  | SIMFCHAR | :CALC:EPAR:SIM
| Statistical analysis | On/Off |  | MEASTAT | :CALC:MARK:MATH:STAT
|  | Read analysis result |  | OUTPMSTAT? | :CALC:MARK:MATH:DATA
| Band Width Analysis | On/Off |  | WIDT | :CALC:MARK:BWID
|  | Cutoff point definition | definition type | WIDVTYPE | :CALC:MARK:BWID:TYPE
|  |  | Fixed value setting | WIDFVAL | :CALC:MARK:BWID:FIX:VAL
|  | Read analysis result |  | OUTPMWID? | :CALC:MARK:BWID:DATA
|  | Inside cutoff point search |  | WIDSIN | Not Available
|  | Outside cutoff point search |  | WIDSOUT | Not Available
| Waveform analysis | Condition | Range setting/Specified range | ANARNG | :CALC:FUNC:DOM
|  |  | Range setting/Sweep range | ANARFULL | :CALC:FUNC:DOM
|  |  | Segment selection | ANASEGM | Not Available
|  |  | Trace selection/Data trace | ANAODATA | Not Available
|  |  | Trace selection/Memory trace | ANAOMEMO | Not Available
|  |  | Peak definition | THRR | :CALC:FUNC:PEXC
|  | Analysis commands | Maximum value search | OUTPMAX? | :CALC:FUNC:TYPE MAX :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Minimum value search | OUTPMIN? | :CALC:FUNC:TYPE MIN :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Maximum and minimum values search | OUTPMINMAX? | Not Available (Use max and min value search)
|  |  | Maximum positive peak search | PEAK? | CALC:MARK:FUNC:PPOL POS :CALC:FUNC:TYPE PEAK :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Next maximum positive peak search | NEXPK? | Not Available (Use the multi peak search)
|  |  | Minimum negative peak search | NPEAK? | CALC:MARK:FUNC:PPOL NEG :CALC:FUNC:TYPE PEAK :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Next minimum negative peak search | NEXNPK? | Not Available (Use the multi peak search)
|  |  | nth positive peak search | LMAXS? | Not Available (Use the multi peak search and find the nth peak)
|  |  | nth negative peak search | LMINS? | Not Available (Use the multi peak search and find the nth peak)
|  |  | Specified measurement parameter value search/To right | TARR? | :CALC:MARK:FUNC:TARG :CALC:FUNC:TYPE RTAR :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Specified measurement parameter value search/To left | TARL? | :CALC:MARK:FUNC:TARG :CALC:FUNC:TYPE LTAR :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Analysis of maximum difference between positive and negative peaks/Value | RPLPP? | :CALC:FUNC:TYPE RPP :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Analysis of maximum difference between positive and negative peaks/Value and peak position | RPLPPS? | Not Available (Use the peak search)
|  |  | Analysis of maximum difference between positive peak and its adjacent negative peak/Both sides | RPLHEI? | :CALC:FUNC:TYPE RPLH :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Analysis of maximum difference between positive peak and its adjacent negative peak/Left side | RPLLHEI? | :CALC:FUNC:TYPE RPLL :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Analysis of maximum difference between positive peak and its adjacent negative peak/Right side | RPLRHEI? | :CALC:FUNC:TYPE RPLR :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Resonator analysis/0-phase point | OUTPRESO? | :CALC:FUNC:TYPE RESO :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Resonator analysis/0-phase point and ripple | OUTPRESR? | :CALC:FUNC:TYPE RESR :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Resonator analysis/Ceramic resonator parameter | OUTPCERR? | :CALC:FUNC:TYPE CERR :CALC:FUNC:EXE :CALC:FUNC:DATA?
|  |  | Resonator analysis/Equivalent circuit parameter of crystal resonator | EQUCPARS4? | :CALC:EPAR:S4?
|  |  | Resonator analysis/parallel capacitance | EQUC０? | :CALC:EPAR:PC0?
Save, recall and so on | Save | State (meas. condition, memory/user cal. data/fixture compen. data array) |  | SAVDSTA | :MMEM:STOR:STYP :MMEM:STOR:STAT
|  | Screen graphics |  | SAVDTIF | :MMEM:STOR:IMAG
|  | Internal data arrays | ASCII file | SAVDASC | Not Available
|  |  | Binary file | SAVDDAT | Not Available
|  |  | Touchstone file (Data array only) | SAVDS1P | :MMEM:STOR:S1P
|  | Select | Data array | SAVDAT | :MMEM:STOR:FDAT
|  |  | Memory array | SAVMEM | :MMEM:STOR:STYP STATe DSTate (Memory trace is also saved into the state file.)
|  |  | Data trace array | SAVDTRC | :MMEM:STOR:STYP STATe DSTate (Data trace is also saved into the state file.)
|  |  | Memory trace array | SAVMTRC | :MMEM:STOR:STYP STATe DSTate (Memory trace is also saved into the state file.)
|  |  | User cal. data / fixture compen. data array | SAVCAL | :MMEM:STOR:STYP STATe CSTate (Calibration is also saved into the state file.)
| Update |  |  | RESAVD | Not Available
| Save state as AUTOREC.STA |  |  | SAVPSTA | :MMEM:STOR "Autorec.sta"
| Recall |  |  | RECD | :MMEM:LOAD
| Storage device selection |  |  | STOD | Use Windows Drive C:/D:/E:
| Directory | Create |  | CRED | :MMEM:MDIR
|  | Move |  | CHAD | Not Available
| File | Copy |  | FILC | :MMEM:COPY
|  | Delete |  | PURG | :MMEM:DEL
| Disk formatting |  |  | INID | Not Available
Status report structure | Clear registers |  |  | CLES | *CLS
| Set instrument event status enable register |  |  | ESNB | Use the qestionable register
| Read instrument event status register |  |  | ESB? | Use the qestionable register
| Set operation status enable register |  |  | OSE | :STAT:OPER:COND
| Set positive transition filter of operation status register |  |  | OSPT | :STAT:OPER:PTR
| Set negative transition filter of operation status register |  |  | OSNT | :STAT:OPER:NTR
| Read operation status event register |  |  | OSER? | :STAT:OPER:EVEN?
| Read operation status condition register |  |  | OSR? | :STAT:OPER:COND
LAN | IP address setting |  |  | ADDRIP | Not Available
| Gateway IP address setting |  |  | ADDRGW | Not Available
| Subnet mask |  |  | SUBNET | Not Available
Printer | Displayed item selection | Trace (Normal display) |  | RESD | Not Available
|  | list of sweep parameter and measurement parameter values | Display | LISV | Not Available
|  |  | Display next page | NEXP | Not Available
|  |  | Display previous page | PREP | Not Available
|  | list of measurement condition parameters |  | OPEP | Not Available
|  | List sweep table | Display | DISL | :DISP:TABL:TYPE SEGM :DISP:TABL:STAT
|  |  | table format selection | DISMPRM | Not Available
|  | Limit line table | Display | DISLLIST | :DISP:TABL:TYPE LIM :DISP:TABL:STAT
|  |  | table format selection | DISMAMP | Not Available
| Printing | Print |  | PRINALL | HCOP:IMM
|  | Cancel |  | COPA | Not Available
|  | Setting | Reset | DFLT | Not Available
|  |  | Resolution | DPI | Not Available
|  |  | Formfeed On/Off | FORMFEED | Not Available
|  |  | Direction | LANDSCAPE | Not Available
|  |  | Left margin | LMARG | Not Available
|  |  | Top margin | TMARG | Not Available
|  |  | Softkey label printing On/Off | PRSOFT | Not Available
|  |  | Time stamp function On/Off | COPT | Not Available
|  |  | Color | PRIC | Not Available
LCD | Screen | Intensity |  | INTE | Not Available
|  | Background intensity |  | BACI | :DISP:COL:BACK
| Display element | Element selection |  | COLO | :DISP:COL:GRAT :DISP:COL:LIM :DISP:COL:RES :DISP:COL:TRAC:DATA :DISP:COL:TRAC:MEM
|  | Brightness |  | CBRI
|  | Chroma |  | COLOR
|  | Hue |  | TINT
| Reset | All setting |  | DEFC | :DISP：COL:RES
|  | Specified |  | RSCO | :DISP：COL:RES
I/O port | 8 bit I/O port | Data output |  | OUT8IO | Not Available
|  | Data input |  | INP8IO? | Not Available
| 24 bit I/O port | Data output | Port A | OUTAIO | :CONT:HAND:A:DATA
|  |  | Port B | OUTBIO | :CONT:HAND:B:DATA
|  |  | Port C | OUTCIO | :CONT:HAND:C:DATA
|  |  | Port D | OUTDIO | :CONT:HAND:D:DATA
|  |  | Port E | OUTEIO | :CONT:HAND:E:DATA
|  |  | Port F | OUTFIO | :CONT:HAND:F:DATA
|  |  | Port G | OUTGIO | :CONT:HAND:G:DATA
|  |  | Port H | OUTHIO | :CONT:HAND:H:DATA
|  | Data input | Port C | OUTPINPCIO? | :CONT:HAND:C:DATA
|  |  | Port D | OUTPINPDIO? | :CONT:HAND:D:DATA
|  |  | Port E | OUTPINPEIO? | :CONT:HAND:E:DATA
|  | Port C setting | Set output port | COUT | :CONT:HAND:C:MODE
|  |  | Set input port | CIN | :CONT:HAND:C:MODE
|  | Port D setting | Set output port | DIN | :CONT:HAND:D:MODE
|  |  | Set input port | DOUT | :CONT:HAND:D:MODE
|  | Check INPUT1 pulse |  | INPT? | Not Available
|  | OUTPUT1 setting | High level / Low level | OUT1{H¦L} | Not Available
|  |  | Operation when INPUT1 detects pulse input | OUT1ENV{H¦L} | Not Available
|  | OUTPUT2 setting | High level / Low level | OUT2{H¦L} | Not Available
|  |  | Operation when INPUT2 detects pulse input | OUT2ENV{H¦L} | Not Available
|  | Positive / Negative logic selection | Positive logic | POSL | Not Available
|  |  | Negative logic | NEGL | Not Available
Beep | On/Off setting of the beep to notify the completion of operation |  |  | BEEPDONE | :SYST:BEEP:COMP:STAT
| On/Off setting of the beep when the limit test result is FAIL |  |  | BEEPFAIL | :SYST:BEEP:WARN:STAT
| On/Off setting of the beep when an error occurred or when an invalid key is pressed |  |  | BEEPWARN | :SYST:BEEP:WARN:STAT
Clock | Date display format setting |  |  | DMODE | Not Available
| Date setting |  |  | SETCDATE | :SYST:DATE
| Time setting |  |  | SETCTIME | :SYST:TIME
Key | Disable front panel key and rotary knob |  |  | DSKY | :SYST:KLOC:KBD
| Enable front panel key and rotary knob |  |  | ENKEY | :SYST:KLOC:KBD
| Send key code |  |  | KEY | Not Available
| Show ON KEY LABELS |  |  | USKEY | Not Available (No i-Basic)
Others | Wait for command operation completion |  |  | *WAI | *WAI
| Read 1 when operation completes |  |  | *OPC? | *OPC?
| Set 1 to OPC bit when operation completes |  |  | *OPC | *OPC
| Controller GPIB address setting |  |  | ADDRCONT | Not Available
| Read product information |  |  | *IDN? | *IDN?
| Read option information |  |  | *OPT? | *OPT?
| Self test operation |  |  | *TST | *TST (Always return 0)

## Functions
Function comparision with 4294A
New
4 channels and 4 traces for quick and multiple parameter measurements under various conditions
Win 7 platform
10.4 inch touch screen
6 USB I/F
Small footprint
Web control
Embedded Online Help
Equation editor
Upper and lower frequency range. 40Hz -> 20Hz, 110MHz -> 120MHz
New LF options and upgradability: 10M, 20M, 30M, 50M. LF options don’t support 42941A/42942A, and have slower measurement speed
FW built-in OSC auto level function (ALC). This function was provided as IBASIC sample in the 4294A
User preset / Factory preset function
Instr Setup key for quick basic setup (support frequency sweep only)
Macro run function
Channel copy and recall function
Two new 4-element equivalent circuit model
Changed:
GPIB command doesn’t have compatibility. E4990A has SCPI and it’s compatible with E4991B about common functions.
Display accumulate function
LF options doesn’t support 42941A/42942A
State file is not compatible
Deleted:
Rear 8 bit I/O port (Use 24 bit I/O)
Stimulus overlapping in Segment Sweep (Use multi-channel)
IBASIC (Use vbscript or install any programing language such as Excel and VB. Macro run function is added to run files *vbs, *exe.)
Floppy disk drive (Use USB memory device)
Function to read bias level using maker (Use DC monitor trace)
Status Reporting System
Calibration or Compensation Completed (Use "measuring".)
Sweep completed (Use "measuring"and STAT:OPER:PTR)
Service Routine waiting or Manual Trigger waiting (Use "Waiting for Trigger" and STAT:OPER:PTR)
Search Failed (Use "Target not found")
Point Measure Completed (Use Wait for trigger)